## Slide 1
Deep CT Recon & Hands-on 3
Ge Wang
AI-based X-ray Imaging System (AXIS) Lab
Rensselaer Polytechnic Institute, Troy, New York, USA
Oct. 5, 2021

## Slide 2: MI Schedule for F21
Tue | Topic | Fri | Topic
08/31 | Introduction | 09/03 | Neuron, Network & Backpropagation
09/07 | No Class | 09/10 | Hands-on 1: Python, Colab, & TensorF
09/14 | Architectures & Training | 09/19 | Hands-on 2: MNIST Classification
09/21 | Image Quality I | 09/24 | CT Physics
09/28 | CT Reconstruction | 10/01 | Deep CT Reconstruction
10/05 | Hands-on 3: CT Networks | 10/08 | More CT Contents
10/12 | MRI Physics | 10/15 | Spin-Echo
10/19 | K-Space Theorem | 10/22 | More MRI Contents
10/26 | Mid Exam | 10/29 | Deep MRI Reconstruction
11/02 | Hands-on 4: MRI Networks | 11/05 | Nuclear Physics
11/09 | Nuclear Imaging | 11/12 | US Imaging
11/16 | Optical Imaging | 11/19 | Multimodality Imaging
11/23 | Hands-on 5: Image Conversion | 11/26 | Thanksgiving
11/30 | Image Quality II | 12/03 | Stability, Interpretability, & Others
12/07 | Overview | 12/10 | Final Exam
Office Hour: Teaching Day 5-6pm: https://rensselaer.webex.com/meet/wangg6

## Slide 3: Weighted Contributions
Pixels contributed to a datum with different weights
Error correction-based updates should use the same weights

## Slide 4: Generalized: Matrix A & Its Transpose AT
Ax=y
A: m x n (2 x 3)
x: n x1 (3 x 1)
ATy’=x’
AT: n x m (3 x 2)
y’: m x 1 (2 x 1)
x in R3
y in R2
y’ in R2
x’ in R3

## Slide 5: Simplified: Matrix A & Its Transpose AT
Ax=y
A: 1 x 2
x: 2 x 1)
ATy’=x’
AT: 2 x 1
y’: 1 x 1
x in R2
y in R
y’ in R
x’ in R2

## Slide 6: Specified: Matrix A & Its Transpose AT
Ax1+bx2=y
A: (a b)
x: (x1 x2) T
ay’=x1’
by’=x2’
AT: (a b)T
y’
x
y
y’
x’

## Slide 7: Heuristically-derived SART Formula
Error
Relative Contribution Weight
Relative Rectifying Factor

## Slide 8: AUTOMAP with Complexity
Zhu B, Liu JZ, Cauley SF, Rosen BR, Rosen MS: Image reconstruction by domain-transform manifold learning. Nature 555:487–492, 2018
k-space Data
Reshape
To a Vector
FC Layers to All Pixels
Image
Complexity
# of Data or Pixels per Row
# of views
Notes: In the literature, only a few deep-learning methods were proposed for reconstructing medical images directly from raw data. AUTOMAP, published last year on Nature, is the first work trying to learn the reconstruction process. N denotes the dimension of reconstructed image. AUTOMAP use 3 fully-connected layers to the the mapping between raw k-space data to a corresponding MRI images. The basic idea of AUTOMAP could be illustrated here. Due to the intrinsic relation between frequency domain and time domain, every pixel in k-space domain is related every pixel in image domain. Fully-connected layer is able to learn this relation. Without a doubt, similar idea could be applied to CT since sinograms and images are also correlated. If we directly apply AUTOMAP for CT applications, the complexity would be . N_v donotes the number of view. There are two problems if we use AUTOMAP for CT: 1. Due to the large computational complexity, AUTOMAP is not able to train normal size CT images on modern GPUs. In the AUTOMAP paper, authors used 128*128 images. 2. FC layers imply that all the pixels in both domain are correlated, but this is not the true for CT. In CT imaging, x-ray are used to take line integral at various angle, so a single point in the sinogram only relate to a line in the field of view.
In the literature, only a few deep-learning methods were proposed for reconstructing medical images directly from raw data. AUTOMAP, published last year on Nature, is the first work trying to learn the reconstruction process. N denotes the dimension of reconstructed image. AUTOMAP use 3 fully-connected layers to the the mapping between raw k-space data to a corresponding MRI images. The basic idea of AUTOMAP could be illustrated here. Due to the intrinsic relation between frequency domain and time domain, every pixel in k-space domain is related every pixel in image domain. Fully-connected layer is able to learn this relation. Without a doubt, similar idea could be applied to CT since sinograms and images are also correlated. If we directly apply AUTOMAP for CT applications, the complexity would be 𝑶(𝑵^𝟑×𝑵_𝒗 ). N_v donotes the number of view. There are two problems if we use AUTOMAP for CT: 1. Due to the large computational complexity, AUTOMAP is not able to train normal size CT images on modern GPUs. In the AUTOMAP paper, authors used 128*128 images. 2. FC layers imply that all the pixels in both domain are correlated, but this is not the true for CT. In CT imaging, x-ray are used to take line integral at various angle, so a single point in the sinogram only relate to a line in the field of view.

## Slide 9: iCTNet with Complexity
*Li, Y., Li, K., Zhang, C., Montoya, J. and Chen, G., “Learning to Reconstruct Computed Tomography (CT) Images Directly from Sinogram Data under A Variety of Data Acquisition Conditions,” IEEE Transactions on Medical Imaging, 1–1 (2019).
sinogram
sinogram
FC
FC
image
image
final output
Notes: The second work trying to learn the reconstruction process is the iCTNet published on IEEE transaction on medical imaging this year. N_c denotes the number of detectors. The overall idea of iCTNet can be summarized here. Assume a CT scanner with 3 detectors and 4-view sinogram is acquired. iCTNet contains a total of 12 layers, in the first 8 layers, convolutional layers are used to de-noised sinogram, and to learn the interpolation process, generating a learned full-view sinogram with number of view , where is a hyperparameter in iCTNet. Then, instead of feeding the whole sinogram into one large FC layer, iCTNet feeds projection at each angle in to a small FC layer, thereby reducing the computation burden. Rotation operations are used to synthesis the analytical FBP method and then one convolutional layer is used to combine all the images at different angle, generating final output.
The second work trying to learn the reconstruction process is the iCTNet published on IEEE transaction on medical imaging this year. N_c denotes the number of detectors. The overall idea of iCTNet can be summarized here. Assume a CT scanner with 3 detectors and 4-view sinogram is acquired. iCTNet contains a total of 12 layers, in the first 8 layers, convolutional layers are used to de-noised sinogram, and to learn the interpolation process, generating a learned full-view sinogram with number of view 4𝛼, where 𝛼 is a hyperparameter in iCTNet. Then, instead of feeding the whole sinogram into one large FC layer, iCTNet feeds projection at each angle in to a small FC layer, thereby reducing the computation burden. Rotation operations are used to synthesis the analytical FBP method and then one convolutional layer is used to combine all the images at different angle, generating final output.

## Slide 10: Our Key Idea: Smart BP with Complexity
Smart Backpropagation with Learned Weights Restricted to a Single X-ray Path
Learned Filtered Projection

## Slide 11: Our Key Idea: Smart BP with Complexity
Smart Backpropagation with Learned Weights Restricted to a Single X-ray Path
Learned Filtered Projection
Notes: Our method is able to learn the reconstruction process as low as O(N) parameters. As illustrated here, blue boxes represent a 3*3 CT image, and 2 projections are acquired at the angles shown here. Yellow dots represent points in the acquired projections. AUTOMAP assumes that these 6 dots are related to all 9 pixels. iCTNet assumes that projection at a particular angle only relates to the back-projected image at this angle. Our method, DNA assumes that every single dot only relates to a single line in the field of view. Which is true for CT scans. And the analytical FBP method also has this assumption. In order words, our method learn the reconstruction process in a point-wise manner. Instead of trying to learn the whole reconstructed image using one large FC layer, proposed method takes a single point as input and tries to learn a single line in the field of view.

## Slide 12: Outline
CT System & Scanning
Cone-beam CT
Spiral Cone-beam CT

## Slide 13: Data Acquisition System (DAS)
Filter
Source
Detector
Pre-Collimator
Post-Collimator
Patient
Scattering

## Slide 14: CT Scanner
Data Acquisition
System
Gantry
Detectors
Storage/Cloud
Display
Control
Console
Computer
Table
Analysis
Source

## Slide 15: CT Scanning Modes

## Slide 16: Outline
CT System & Scanning
Cone-beam CT
Spiral Cone-beam CT

## Slide 17: Equi-spatial Fan-beam Formula

## Slide 18: d(x, y, z) to d(u,v) with Magnification & Distortion
A cone-beam projection of a 3D delta function is a magnified/demagnified 2D delta function with or without distortion

## Slide 19: Obliqueness of the Detector Array

## Slide 20: Divergence of the X-ray Beam
This Scaling Factor Applies the Same Way from the Opposite View Angles (Generally Speaking, Either Magnified or De-magnified Depending on Which Side)

## Slide 21: Difference Between θ & β Angles
For Simplicity, Let Us Only Consider the Impulse (Red) Near the System Origin

## Slide 22: From Full-scan to Half-scan
Full-scan Fan-beam Dataset (360˚) Consists of Two Parallel-beam Datasets (180˚)

## Slide 23: Cone-beam Geometry
Cone-beam geometry is a natural extension of fan-beam geometry
An order of magnitude improvement of data acquisition speed relative to fan-beam geometry
Detector Plate
Source
Object
Feldkamp LA, Davis LC, Kress JW: Practical cone-beam algorithm.
Journal of the Optical Society of America A 1:612-619, 1984

## Slide 24: Fan-beam to Cone-beam Projection Data

## Slide 25: Cone-beam Data Scaled Back
The detector plane mapped onto
the t-z plane
(which is convenient, being consistent to the fan-beam case)

## Slide 26: Cone-Beam to Fan-Beam Data (Point-wise)

X-ray source
Virtual in-plane
X-ray source
Object point
Tilted fan
Tilting angle

## Slide 27: Cone-Beam to Fan-Beam Data (Segment-wise)
Wang G, Zhao SY, Cheng PC: Exact and approximate cone-beam X-ray microtomography (Chap. 1). Focus on multidimensional microscopy (I) (Editors: Cheng PC, Huang PP, Wu JL, Wang G, Kim HG), World Scientific, Singapore, 233-261, 1999

## Slide 28: Cone-beam Data Correction
β
α
General Scanning Trajectory

## Slide 29: Obliqueness Correction
Obliqueness Correction in Fan-beam Formula
Obliqueness Correction in Cone-beam Formula
An X-ray Hits the 2D Detector Plate Obliquely, Determined by Two Angles

## Slide 30: Homework
Read papers #4-6 to understand what I explained in the class on iterative and deep reconstruction
Watch the hands-on3 video, do hands-on 3 assignments, and report your experience in deep denoising and deep reconstruction (Network principle, training & testing time, error analysis)

## Slide 31: Homework
Please run through all these codes according to the explained steps
Report on your experience in using the post-processing and direct mapping methods, principle summary, training and testing time, error relative to the gold standard.