## Slide 1
CT Reconstruction – Fan-beam Geometry
Ge Wang
AI-based X-ray Imaging System (AXIS) Lab
Rensselaer Polytechnic Institute, Troy, New York, USA
September 8, 2023

## Slide 2: Outline
X-ray Nature
Wave
Particle
X-ray Attenuation
Mechanisms
  Beer’s Law
X-ray Tubes/Sources & Detectors
Sources & Detectors
X-ray Radiography
Radiation Dose
Dosimetry
Low-dose CT

## Slide 3: System of Linear Equations

## Slide 4: Derivation of Beer’s Law

## Slide 5
Beer’s Law Gives a Linear Equation
Each beam measures a line integral
Beer’s Law

## Slide 6: Pixels Are Picture Elements

## Slide 7: Sum Game
=4
2=3
3=2
4=1
Practical Problems Involve Many More Unknowns

## Slide 8: Parallel-beam Scanning
p(t)
f(x,y)
Computed tomography (CT): Image reconstruction from projections
p(t)
f(x,y)
t

y
x
X-rays
Radon Transform
t

## Slide 9: Sinusoidal Waves Are Building Blocks

## Slide 10: Fourier Slice Theorem
v
u
θ
F(u,v)
p(θ,t)
f(x,y)
t
θ
y
x
X-rays
F[p(θ,t)]
No information in line integrals along non-vertical directions
Information on horizontal waves are in vertical projection data

## Slide 11: Outline
Parallel-beam Case
Fan-beam Case
Heuristic Extension

## Slide 13: Fourier Slice Theorem

## Slide 15: Parallel-beam FBP

## Slide 16: Equi-angular/spatial Fan-beam Geometry
Homework on Half-scan

## Slide 17: Parallel- to Fan-beam FBP
Parallel-beam
Filtered Backprojection
Fan-beam
Filtered Backprojection
Key Idea: Convert Parallel-beam Data to Fan-beam Data

## Slide 18: Equi-spatial Fan-beam Formula

## Slide 19: Fan-beam FBP = Parallel-beam FBP
Key Idea: Convert Parallel-beam Data to Fan-beam Data

## Slide 20: Fan- to Parallel-beam FBP
Parallel-beam
Filtered Backprojection
Fan-beam
Filtered Backprojection
Key Idea: Convert Fan-beam Data to Parallel-beam Data or We Do It Heuristically!

## Slide 21: Heuristics behind the Fan-beam Formula
If you want to see directly the correctness of the fan-beam formula, you can read the rest of the slides.
The key idea is that if we just consider a point object near the system origin, then all the factors in the fan-beam formula becomes immediately clear!
And that understanding is valid no matter the source to the original distance is large or small. Hence, we got a generalized fan-beam formula!

## Slide 22: Equi-spatial Fan-beam Formula

## Slide 23: Delta Function on a Scaled Axis
Projecting with a Divergent Beam: The Impulse Becomes Stronger (Larger Area)

## Slide 24: Delta Function on a Tilted Axis
Projecting onto a Tilted Plane: The Impulse Becomes Stronger (Larger Area)

## Slide 25: Obliqueness of the Detector Array

## Slide 26: Divergence of the X-ray Beam
This Scaling Factor Applies the Same Way from the Opposite View Angles (Generally Speaking, Either Magnified or De-magnified Depending on Which Side)

## Slide 27: Difference Between θ & β Angles
For Simplicity, Let Us Only Consider the Impulse (Red) Near the System Origin
Notes: Note that dtheta and dbeta are both angular weights. Their sums are both 2pi. Hence, we can make weights for dtheta uniform just as the weights for dbeta is naturally uniform. Hence, the results are the same.
Note that dt’ is for equi-spatial fan-beam data, and needs to be converted to parallel-beam dt. Here we have a cosine factor. Also, dbeta and dtheta should be (U/R)*the same cosine factor

## Slide 28: From Full-scan to Half-scan
Full-scan Fan-beam Dataset (360˚) Consists of Two Parallel-beam Datasets (180˚)

## Slide 29: Equi-spatial Fan-beam Formula

## Slide 30: Extended Fan-beam Formula

## Slide 31: More General Scanning Trajectory

## Slide 32: Simulation Results
Effectiveness of the derivative-free noncircular fan-beam reconstruction (2562 pixels, 256 projections, 256 detectors)
Circle-like and triangle-like random scanning loci
True slice from the 3D Shepp-Logan phantom
Computed using the circle-like locus
Computed using the triangle-like locus

## Slide 33: Homework
What is the angular range [0, x] for β so that fan-beam projection data meet the sufficient condition for image reconstruction? In other words, x=?

## Slide 34: Homework
There is a mathematical error in my paper, please find it, and email me (I am curious to see who is the first)
https://ieeexplore.ieee.org/document/242365/